# ADS1261 Debugging

### Step 1: Test the Board

First, connect all five load cells and run the firmware that continuously reads all five channels.

### Expected Result

- All five channels should give stable readings.

- The readings should change correctly when a load is applied.

### If everything works

The board is working correctly.

### If there is a problem

There are two possibilities:

- Only one or a few channels are not working → Go to Step 2.

- None of the channels are working → Go to Step 4.

## Step 2: Check the Faulty Channel

Replace the load cell connected to the faulty channel with a known good load cell.

- If the problem moves with the load cell, the load cell is faulty.

- If the same channel is still faulty, the problem is on the ADC board.

Next, connect that same load cell to a known working board.

- If it works on the good board, the ADC board is confirmed to be faulty.

## Step 3: Check the Analog Input Voltages

Measure the voltages at AINP and AINN of the faulty channel.

With a 5 V excitation and no load, both voltages should be around 2.5 V.

For example:

Working board

- AINP ≈ 2.5 V

- AINN ≈ 2.5 V

Faulty board

- AINP ≈ 1.5 V

- AINN ≈ 2.5 V

If the two voltages are not approximately equal, the analog channel is faulty.

## Step 4: Check SPI Communication

If all five channels are not working, run the Chip ID test program.

### Expected Result

The ADS1261 Chip ID should be detected correctly.

### If the Chip ID is detected

SPI communication is working.

Now run the Write Register / Read Register test.

- If register write and read are successful, the SPI interface and ADS1261 are working correctly.

- The problem is likely in the analog input circuitry or channel hardware.

### If the Chip ID is not detected

The problem is likely related to SPI communication, power, or hardware.

Proceed to the next step.

## Step 5: Check Power Supply and Connections

Verify:

- DVDD = 3.3 V

- AVDD = 5 V

If either voltage is lower than expected, there may be a power supply issue, short circuit, or damaged ADS1261.

Also check the continuity of:

- CS

- SCLK

- MOSI

- MISO

- DRDY

- RESET

- Power and Ground

## Step 6: Check the DRDY Signal

Observe the DRDY pin.

A working board should generate regular active-low pulses.

Possible faults:

- No DRDY pulse

- Random pulses

- Positive pulses instead of active-low pulses

These indicate that the ADS1261 is not operating correctly.

## Step 7: Compare with a Working Board

Finally, compare the faulty board with a known good board using the same:

- Firmware

- Teensy

- Load cells

- Power supply

Compare:

- DVDD and AVDD

- AINP and AINN voltages

- Chip ID

- Register read/write

- DRDY signal

- ADC readings

This makes it much easier to identify whether the issue is with the PCB, assembly, power supply, SPI communication, or the ADS1261 itself.